[image: image4.png]Agilent EducatorsCorner.com Experiments
e ———————————————— —





[image: image4.png]
Digitizing Oscilloscope Workshop

Triggering & Observing & Measuring A Complex Waveform

By:
Walter Banzhaf, P.E.


Professor of Engineering Technology


University of Hartford


Ward College of Technology

Purpose:

Oscilloscopes can be rather obtuse creatures: they trigger just as you tell them to (source, slope, and level specified), and they don’t “look” at the display to see if it’s correct. For example, in Figure 1 the display is obviously not giving a correct picture of the waveform, since it shows two pulses as having both low and high levels present simultaneously. This procedure will acquaint you with sweep holdoff, which can be used to correct this problem. Sweep holdoff delays the next sweep until the user-controlled holdoff time has expired and a subsequent trigger has occurred; this avoids triggering on different pulses in an irregular pulse train.  Unlike most analog ‘scopes, the holdoff time on the digitizing oscilloscope doesn’t change as the time/division is changed. This means that once you’ve adjusted the holdoff time correctly, you can change the time/div.

Equipment Required:


•
Agilent 54600 - Series Oscilloscope

[This experiment is written for the 4-channel Agilent 54601B - but could be easily adapted to a 2-Channel scope].

Introduction 

Circuit Explanation:  

The digital waveform generator (DWG) used in the digital fundamentals laboratory course will be used to produce an irregular pulse train.  It needs only to be connected to a +5 V power supply.

Procedure A - Observing An Improperly Triggered Waveform:
1)
Connect the DWG to a 5 V power supply, and turn on power.  Observe DWG output G on channel 1, using a 10X probe.  Press  SETUP, then press Default Setup.  Set the probe type for channel 1 to 10X by pressing 1 and Probe.  Finally, press AUTO-SCALE. The display should look something like Figure 1.  There’s a triggering problem, since the 2nd and 3rd pulses have both low and high levels present simultaneously.

Procedure B - Using Holdoff To Properly Trigger On The Waveform:
1)
While still looking at the G waveform, rotate the Holdoff control, located in the TRIGGER section, until the displayed waveform is “correct”.  Adjust the time/div as needed in order to see the waveform.  The true pattern of pulses can now be seen clearly, with each pulse having a high level only.

Procedure C - Observing Four Channels Simultaneously:
1)
Connect Channels 1, 2, 3 and 4 to DWG outputs A, B, C and D, respectively using 10X probes (set the probe type for each channel to 10X).

2)
Press  AUTO-SCALE ; the display should look like Figure 3.

3)
Measure the frequency of Channels 1, 2, 3 and 4 using the Measure Time hardkey and the Frequency softkey.

4)
Try measuring the time delay between the rising edge of Channels 3 & 4 using the Measure Cursors hardkey and the t1 & t2 softkeys.

5)
Press the Source hardkey in the Trigger section; notice that the Autoscale function had selected Channel 4 as the trigger source.  It is just dumb luck that Channel 4 had the lowest frequency of the four channels, and as a result a stable display occurred.

6)
Rotate inputs as follows: Channels 1 and 4 to DWG outputs D and A, respectively.  Press AUTO-SCALE ; the display is no longer stable, as Channel 4 was selected again as trigger source, but Channel 4 is connected to the highest frequency in the system.

[image: image1.png]<—0.00 s 1.00%/

1HH

r
'

THH 2 HH

Vmax(1)=4.000 V Vmin(1)=31.25mV

Period(1)=8.305ms

State  Volts/Div Position Couplg BW Lim Invert Prove
Chanl On 1.000 V -2.000 V DC Off Off 10:1
Chan 2  Off 1.000 V 0.000V DC Off Off 10:1
Chan 3  Off 100.0mV 0.000V DC --- --- 1:1
Chan 4  Off 100.0mV 0.000 V DC - - 1:1
Main Main Time Delayed Delayed
Mode Time/Div Delay Ref Time/Div  Delay
Horizontal Normal 1.000ms/ 0.000 s Left
Trigger Mode Source Level Holdoff Slope Couplg Reject NoiseRe;
AutoLvl Ch1l 1.281mV 200.0ns Pos DC Off off

Display Mode: Normal

Figure 1 - Irregular Pulse Waveform (G of DWG), Improperly Triggered




[image: image2.png]1 1.00V

41 srop

1444 : : N L E

: : : ; : : 1
..... e S S

' ' . ' f ' ' . '
State  Volts/Div Position Couplg BW Lim Invert Prove
Chanl On 1.000 V -1.094V DC Off Off 10:1
Chan 2  Off 1.000 V 0.000V DC Off Off 1:1
Chan 3  Off 100.0mV 0.000 V DC -- - 1:1
Chan 4  Off 100.0mV 0.000 V DC -—- - 1:1
Main Main Time Delayed Delayed
Mode Time/Div Delay Ref Time/Div  Delay

Horizontal Normal 2.000ms/

Trigger Mode Source Level

0.000 s
Holdoff Slope

Cntr

AutoLvl Chl 1.031mV 6.680ms Pos

Display Mode: Normal

Couplg
DC

Reject NoiseRej
Off Off

Figure 2 - Irregular Pulse Waveform (G of DWG),
Properly Triggered Using Holdoff




[image: image3.png]0.00s 2.007/ PN 44 stop

1 5.00V 2 5.00V 35V 45V r
< et
= 1
- 2
- : - : = : - : : 1
IR R e N R R R R N R R R E NN R EEEE
. . = Z 3
I S\
== I T - - R [ i i R " e D DN 4

State  Volts/Div Position Couplg BW Lim Invert Prove

Chan 1 On 5.000 V 12.50 V DC Off Off 10:1
Chan2 On 5.000 V 3.000 V DC Off Off 10:1
Chan3 On 5.000 V -7.219V DC --- - 10:1
Chan4 On 5.000 V -16.72 V DC - --- 10:1
Main Main Time Delayed Delayed
Mode Time/Div Delay Ref Time/Div  Delay

Horizontal Normal 2.000ms/ 0.000s Cntr

Trigger Mode Source Level Holdoff Slope Couplg Reject NoiseRe]
AutoLvl Ch4 1.719V  200.0ns Pos DC Off Off

Display Mode: Average # Average: 8

Figure 3 - Waveforms A, B, C and D of DWG,
Properly Triggered on Channel 4




� EMBED Word.Picture.8  ���








[image: image5.wmf]


_1030875879.doc
[image: image1.png]Agilent Technologies
EducatorsCorner.com Experiments








