[image: image7.png]Agilent EducatorsCorner.com Experiments
e ———————————————— —





[image: image7.png]
Turn-on & Turn-off Times of a Power Supply

By:
 Walter Banzhaf

P.E. Professor of Engineering Technology

University of Hartford

Ward College of Technology
Abstract:

Regulated power supplies are found in electronics laboratories, and in virtually every electronic circuit. They are either switching power supplies, which are quite efficient but rather complex and prone to generate EMI  (electro-magnetic interference), or linear power supplies, which are simpler but often highly inefficient. One parameter of a power supply is how quickly it comes on line after being turned on. This dV/dt will affect the current that flows into a capacitive (or tungsten filament) load, and may determine whether certain types of oscillators will start. Most oscillators will start faster, and are more likely to begin sustained oscillations, if the power to which they are connected is turned on quickly.

Equipment:

· Agilent 54600 - Series Oscilloscope 

· Any 2-Output Power Supply 

Circuit Explanation:

The regulated dual power supply shown below will be monitored. By putting the oscilloscope in single-shot mode, and by triggering on the rise of voltage on Channel 1, the turn-on dV/dt of both the positive and negative adjustable supplies can be captured when the push button switch is pushed to ON.

[image: image1.png]+ =20 CHANNEL 1

0to 20 V_=
BLK

+

0to20V =—
- GRN ~ G HANNEL 2 |




               Figure 1

Procedure A - Observing And Measuring ± 20 V Supplies:

1) Set (by eye is fine) the PS 503A variable supplies for about ±16 V. Turn both current limit controls fully CCW (minimum current). Connect Channels 1 and 2, using 10X probes, as shown.

2) Refer to the information in Figure 2 for oscilloscope control settings, and adjust your oscilloscope accordingly. The trigger level is not critical, but should be somewhere around +1 V or +2 V.

3) Turn on the power supply output by pushing the ON button IN. The sweep of the oscilloscope should be triggered once, and a new trace recorded (see Figure 2 for a typical display). You will have to press the Run hardkey after each trigger to "arm" the sweep again.

4) Press the Stop hardkey, then press the Display hardkey followed by the Vectors On softkey. Vectors On essentially "connects the dots", giving a better display of the trace for Channel 2. Notable is the pre-trigger information to the left of the trigger time. In other words, even though the oscilloscope had not yet triggered as the rising voltage of the positive power supply neared the trigger voltage, the analog-to-digital converter in the oscilloscope was sampling the power supply output continuously and feeding it into its memory. When the trigger finally occurred, the "sweep" finished and the contents of the memory (containing both pre- and post-trigger information) were frozen. Record turn-on times in the table below.

5) Press the Run hardkey, set the trigger level to about +15 V, and change the trigger slope to negative. See Figure 3. Turn off the power supply output by pushing the ON button again so that it comes OUT. The sweep of the oscilloscope should be triggered once, and a new trace recorded (see Figure 2 for a typical display). Record turn-off times in the table below.

	Power Supply
	Turn-on Time
	Turn-off Time

	+20V (Red)
	
	

	- 20V (Green)
	
	


Procedure B - Observing And Measuring + 5 V Supply:

1) Connect Channel 1 to the +5 V supply. Refer to the information in Figure 4 for oscilloscope control settings, and adjust your oscilloscope accordingly. The trigger level is not critical, but should be somewhere around +0.5 V.

[image: image2.png]== CHANNEL 1





2)Turn on the power supply output by pushing the ON button IN. The sweep of the oscilloscope should be triggered once, and a new trace recorded (see Figure 4 for a typical display). You will have to press the Run hardkey after each trigger to "arm" the sweep again.

3) Press the Stop hardkey. Notable is the much shorter (by three orders of magnitude!) turn-on time, and the non-monatomic rise from 0V to +5 V. Record turn-on time in the table below.

4) Press the Run hardkey, set time/div to 5 ms/div, adjust the trigger level to about +4 V, and change the trigger slope to negative. See Figure 5. Turn off the power supply output by pushing the ON button again so that it comes OUT. The sweep of the oscilloscope should be triggered once, and a new trace recorded (see Figure 5 for a typical display). Record turn-off time in the table below.

	Power Supply
	Turn-on Time
	Turn-off Time

	5V
	
	


[image: image3.png]15600V 25.00V +0.00s 10037 Sngl 3 1 STOP
v T

Frisl 44

00ms t2 = 77.40ms At = 80.80ms 1/At = 12.38 Hz

State  Volts/Div  Position Couplg BW Lim Invert Prove

Chanl On 5.000 V 0.000 V DC off Off 10:1
Chan2 On 5.000 V. 0.000 V DC off off 10:1
Chan3  Off 100.0mV 0.000 V DC - 1:1
Chan4 Off 100.0mV 0.000 V DC - L1

Main Main Time Delayed Delayed
Mode  Time/Div Delay  Ref  Time/Div Delay
Horizontal Normal 10.00ms/ 0.000s  Left

Trigger Mode Source Level Holdoff Slope Couplg Reject NoiseRej
Single Ch1 1.250 V. 200.0ns Pos DC off off

Display Mode: Normal
Cursors: t1=-3.400ms t2=77.40ms VI(1)=

000V V2(1)=0.000 V

Figure 2 - * 20 V Supplies, After Turn-on





[image: image4.png]15.00V__ 25.00V
T v

1/At = 18.32 Hz

State  Volts/Div  Position ~ Couplg BW Lim Invert Prove
Chanl On 5.000 V 0.000 V DC ofe ot 10:1
Chan2 On 5.000 V 0.000 V DC off off 10:1
Chan3 Off 100.0mV 0.000 V DC - 1:1
Chan4  Off 100.0mV 0.000 V. DC

Main Main Time Delayed Delayed
Mode  Time/Div Delay Ref  Time/Div Delay
Horizontal Normal 10.00ms/ -10.000ms Left e oee

Trigger Mode Source Level Holdoff Slope Couplg Reject NoiseRej
Single Ch1 15.00V 200.0ns Neg DC off off

Display Mode: Normal
Cursors: t1=-8.600ms t2=46.00ms V1(1)=562.5mV  V2(1)=5.000 V

Figure 3 - *20 V Supplies, After Turn-off




[image: image5.png]10.0:/ 41_sTop

L--3..

Vmin(1)=31.25mV

Vmax(1)=5.219 V Vtop(1)=5.031 V

State  Volts/Div  Position  Couplg BW Lim Invert Prove

Chan1 On 1.000 V 3000V DC off off 10:1

Chan2 Off  5.000V 0.000 V DC off off 10:1

Chan3 Off  100.0mV  0.000V DC - 1:1

Chan4 Off  100.0mV  0.000V DC - L1
Main Main Time Delayed Delayed

Mode Time/Div Delay Ref Time/Div  Delay
Horizontal Normal 10.00us/ 0.000s Left  —oeeeeeees o

Trigger Mode Source Level Holdoff Slope Couplg Reject NoiseRe;
Single Ch 1 468.8mV 200.0ns Pos DC Off Off

Display Mode: Normal

Figure 4 - +5 V Suppy, After Turn-on




[image: image6.png]1/At = 61.35 Hz

State Volts/Div  Position  Couplg Invert  Prove

Chanl On 1000V -3.000V  DC ofr 10:1
Chan2 Off 5000V 0000V  DC of 10:1
Chan3 Off  1000mV  0.000V  DC 1:1

Chan4 Off  1000mV 0000V  DC

Main Main Time Delayed Delayed
Mode  Time/Div Delay ~ Ref  Time/Div Delay
Horizontal Normal 5.000ms/ 0.000s  Left -

Trigger Mode Source Level  Holdoff Slope Couplg Reject NoiseRej
Single Ch1  4.094mV 2000ns Neg DC off  of

Display Mode: Normal
Cursors: t1=-1.300ms t2=15.00ms VI(1)=662.5mV  V2(1)=5.000 V

11

Figure 5 - +5 V Suppy, After Turn-off




� EMBED Word.Picture.8  ���











[image: image8.wmf]


_1030875879.doc
[image: image1.png]Agilent Technologies
EducatorsCorner.com Experiments








