[image: image8.png]Agilent EducatorsCorner.com Experiments
e ———————————————— —





[image: image8.png]
Curve Tracer

By:  Michael A. Covington


Artificial Intelligence Center

The University of Georgia

Athens, GA  30602-7415

mc@uga.edu
Equipment

•
 Agilent E3631A Power Supply


•
 Computer


•
 RS-232 module

Technique
A graph of current versus voltage (an I/V curve) is a basic characteristic

of almost any electronic device. Together with a PC, the Agilent E3631A power

supply can measure this curve for a wide variety of devices and even capture

full families of curves for power transistors.

Here's an example, the linear I/V curve of a 47-ohm resistor:

[image: image1.png]47-ohn resistor
810 12
81 .3





And here's the I/V curve of a red LED, which "turns on" at 1.6 volts:

[image: image2.png]Red LED
81 5
810 .015





These curves were captured with Program 1, shown at the end of this document. The program incorporates a couple of subtle tricks. First, since small currents are often of interest, but the measurement error of the E3631A can be up to 0.5 mA, the program determines the offset error (by applying 0 volts and measuring the current), then subtracts it from the measured values. As a result, small currents are measured more accurately, although the resolution is still only 1 mA. Second, to avoid excessive current, the program terminates each trace as soon as it goes off the top of the screen.

Thermal Effects
It should be emphasized that the voltages and currents supplied and measured by the E3631A are suitable for relatively high-power devices; this is not the best way to test a single 1N4148 or the like. However, the testing of higher-power components can be quite revealing. Consider this I/V curve of a No. 47 lamp (a 6.3-volt, 150-mA light bulb familiar to radio repairmen of the 1950s):

[image: image3.png]#47 bulb
81 7
810 .25





As expected, the curve is concave downward, but what's the strange hook at the left? Apparently, it reflects the fact that when its heat output is relatively low, the filament doesn't have time to warm up before the measurement is taken. Here's the same curve, but this time with a 5-second delay before taking each measurement:

[image: image4.png]#47 bulb, 5 sec delay betueen measurenents
81 7
810 .25

v
1
[ ]





This time the curve is much smoother but continues to be concave downward, showing that the resistance of the bulb rises with the voltage. This nonlinearity was used in Agilent Technologies’ very first product, the Model 200 audio oscillator, which used an incandescent bulb to stabilize a vacuum-tube Wien bridge.

Transistor Curve Tracing
The E3631A can capture the family of curves of a power transistor by plotting I against V in the manner just described while applying the base current from another of its voltage outputs.

The catch is that typically, the base current needs a resolution much better than 1 mA; it's typical to step from 0 to 1 mA in five steps. Thus the base current cannot be taken directly from an output of the E3631A in current-limiting mode.

A more conventional engineer might have gotten around this by building a voltage-controlled current source (VCCS), but I wanted to keep the circuit as simple as possible, so I used the simplest VCCS possible, a resistor. Here's the complete transistor-curve-tracing circuit:

[image: image9.wmf][image: image5.png]HP
E3631A





To get 1 mA to flow into the base, the computer applies a voltage of (4700 ohms x 1 mA) + 0.6, where 0.6 V is the "turn-on" voltage for a silicon junction. But when trying out the circuit, I quickly discovered that the "turn-on" voltage varies from transistor to transistor. Because of this, instead of using a fixed value of 0.6 V, the program actually measures the turn-on voltage at which current first begins to flow. By doing this, it easily achieves a resolution of 0.02 mA of base current.

The program for transistor curve tracing is Program 2 at the end of this document. Here's its family of curves for a TIP33 power transistor:

[image: image6.png]TIP33 (NPN Si)
vy =8 10 5

IC) =8 T0 .15

IB) =8 10 .081
Measuring U(BE)... .62





And here are the curves for the same transistor, measured on a Tektronix 575 transistor curve tracer (vintage 1957) and photographed with a digital camera:

[image: image7.png]JII]II]IHII




Although no substitute for the Tek 575, the E3631A, with computer control, is in some ways more versatile; besides testing transistors, it's easy to set up simple curve-tracing test jigs for FETs, thyristors, relays (measuring pull-in and drop-out voltage), and a wide variety of other components. The biggest drawback of the E3631A is that it takes a minute or more to measure each trace; with power devices, where heating is a factor, the time delay can actually make the trace more informative.

Program 1 (HPCURVE.BAS)
CLS

PRINT "Current/voltage curve of any component - M. Covington 1999"

PRINT

PRINT "Connect component under test to +25V output of Agilent E3631A."

PRINT

PRINT "Establishing communications with power supply..."

OPEN "COM2:9600,n,8,2,rs,cd,lf,pe" FOR RANDOM AS #1 LEN = 1000

PRINT #1, "SYST:REM"
' Place E3631A in remote mode

PRINT #1, "*RST;*CLS"
' Reset it and clear errors

PRINT #1, "*IDN?"
' Request its ID string

LINE INPUT #1, r$

PRINT r$

PRINT

PRINT "Title of chart?"

LINE INPUT title$

INPUT "V(MAX)?", vmax

INPUT "I(MAX)?", imax

steps = 100
' determines smoothness of graph

' Go into graphics mode

SCREEN 12

xleft = 10: xright = 630: ytop = 10: ybottom = 470

LINE (xleft, ytop)-(xright, ybottom), 4, B     ' Draw the box

PRINT

PRINT " "; title$

PRINT " V = 0 TO "; vmax

PRINT " I = 0 TO "; imax

' For greater accuracy, measure the

' zeroing error in the E3631A's current reading

PRINT #1, "APPL P25V,0,0.01"

PRINT #1, "OUTP ON"

PRINT #1, "MEAS:CURR? P25V"

LINE INPUT #1, i$

ioffset = VAL(i$)

' Ramp up the voltage, measuring the current along the way

FOR vcc = 0 TO vmax STEP vmax / steps

  PRINT #1, "APPL P25V,"; vcc; ","; imax + .001

  ' Time delay for stabilization --

  ' Needed only when heating of DUT is a factor

  'SLEEP 5

  PRINT #1, "MEAS:VOLT? P25V"

  LINE INPUT #1, v$

  v = VAL(v$)

  PRINT #1, "MEAS:CURR? P25V"

  LINE INPUT #1, i$

  i = VAL(i$) - ioffset

  x = xleft + (v / vmax) * (xright - xleft)

  y = ybottom + (i / imax) * (ytop - ybottom)

  IF vcc = 0 THEN PSET (x, y) ELSE LINE -(x, y), 7

  IF i > imax THEN GOTO EndOfThisCurve  ' avoid overcurrent

NEXT vcc

EndOfThisCurve:

PRINT #1, "OUTP OFF"

CLOSE #1

' Wait for user to press Enter

' (Cursor becomes visible when this happens)

PRINT " ";

LINE INPUT junk$

END

Program 2 (HPTRANS.BAS)
CLS

PRINT "TRANSISTOR CURVE TRACER - M. Covington 1999"

PRINT

PRINT "Tests NPN transistors using Agilent E3631A."

PRINT "Base to +6V through 4.7K resistor."

PRINT "Collector to +25V."

PRINT "Emitter to ground of both supplies."

PRINT

PRINT "Establishing communications with power supply..."

OPEN "COM2:9600,n,8,2,rs,cd,lf,pe" FOR RANDOM AS #1 LEN = 1000

PRINT #1, "SYST:REM"

PRINT #1, "*RST;*CLS"

PRINT #1, "*IDN?"

LINE INPUT #1, r$

PRINT r$

PRINT

PRINT "Title of chart?"

LINE INPUT title$

INPUT "V(CC)(MAX) (volts)? ", vccmax

INPUT "I(C)(MAX)  (amps) ? ", icmax

INPUT "I(B)(MAX)  (amps) ? ", ibmax

steps = 100

curves = 5

resistor = 4700

' Go into graphics mode

SCREEN 12

xleft = 10: xright = 630: ytop = 10: ybottom = 470

LINE (xleft, ytop)-(xright, ybottom), 4, B   ' Draw the box

PRINT

PRINT " "; title$

PRINT " V(C) = 0 TO "; vccmax

PRINT " I(C) = 0 TO "; icmax

PRINT " I(B) = 0 TO "; ibmax

' For greater accuracy, measure the

' zeroing error in the E3631A's current reading

PRINT #1, "APPL P25V,0,0.01"

PRINT #1, "OUTP ON"

PRINT #1, "MEAS:CURR? P25V"

LINE INPUT #1, i$

ioffset = VAL(i$)

' Now determine the lowest base voltage that will cause

' detectable current to flow...

PRINT " Measuring V(BE)... ";

i = 0

v = .1

PRINT #1, "APPL P25V,"; vccmax / 2; ",0.01"

WHILE i < ioffset + .001

  v = v + .02

  PRINT #1, "APPL P6V,"; v; ","; ibmax + .005

  PRINT #1, "MEAS:CURR? P25V"

  LINE INPUT #1, i$

  i = VAL(i$)

WEND

vbe = v - .02  ' go back one step

PRINT vbe

' Family of curves...

FOR ib = 0 TO ibmax STEP ibmax / curves

  vb = ib * resistor + vbe

  PRINT #1, "APPL P6V,"; vb; ","; ibmax + .005

  IF vb > 6.1 THEN PRINT "Caution: You need a smaller resistor."

  FOR vcc = 0 TO vccmax STEP vccmax / steps

    PRINT #1, "APPL P25V,"; vcc; ","; icmax + .001

    ' Time delay for stabilization --

    ' Needed only when heating of DUT is a factor

    'SLEEP 5

    PRINT #1, "MEAS:VOLT? P25V"

    LINE INPUT #1, v$

    v = VAL(v$)

    PRINT #1, "MEAS:CURR? P25V"

    LINE INPUT #1, i$

    i = VAL(i$) - ioffset

    x = xleft + (v / vccmax) * (xright - xleft)

    y = ybottom + (i / icmax) * (ytop - ybottom)

    IF vcc = 0 THEN PSET (x, y) ELSE LINE -(x, y), 7

    IF i > icmax THEN GOTO EndOfThisCurve  ' avoid overcurrent

  NEXT vcc

EndOfThisCurve:

NEXT ib

PRINT #1, "OUTP OFF"

CLOSE #1

' Wait for user to press Enter

' (Cursor becomes visible when this happens)

PRINT " ";

LINE INPUT junk$

END

� EMBED Word.Picture.8  ���











_1030875879.doc
[image: image1.png]Agilent Technologies
EducatorsCorner.com Experiments








